# Summary Results from the UK NO2 Network Field Intercomparison Exercise 1999

## 1. Introduction

### 1.1 Background

Within the DETR's UK NO2 Network, there are currently 33 analytical laboratories responsible for preparation and analysis of the NO2 diffusion tubes used by Local and Unitary Authorities. It is important that the data obtained by this Network is of the highest possible quality, therefore QA/QC of laboratory analyses is very important to the Network's operation.

Since 1996, the UK NO2 Network has employed a laboratory performance testing programme, which uses artificially doped diffusion tubes to test the accuracy of laboratory analyses on a monthly basis (Bush et al)1. This scheme provides valuable information on the analytical performance of laboratories. However, owing to the use of artificially doped diffusion tubes in this programme, uncertainties arising from the sampling phase of diffusion tube monitoring cannot be assessed. Field intercomparison exercises were therefore reintroduced into the Network’s laboratory QA/QC programme, in order to better define the sampling and analytical performance of diffusion tubes under normal operating conditions.

The 1999 field intercomparison exercise follows on from previous intercomparisons held in 1993, 1994, 1995 (Stevenson et al)2,3,4 and most recently 1998(Bush) 5. It was designed to complement the existing laboratory quality assurance programme for the UK NO2 Network, which currently utilises information supplied by the WASP scheme under the management of the Health and Safety Laboratory (HSL).

Full details of the performance of individual laboratories in the WASP scheme and the 1999 field intercomparison exercise are available direct from the laboratory.

### 1.2 Conclusions drawn from the 1998 Field Intercomparison

In 1998, a three-month field intercomparison exercise was carried out, involving the thirty-eight laboratories participating at that time. The main conclusions that may be drawn from the 1998 field intercomparison exercise are as follows:

- 79% of the 38 laboratories performed to the data quality objectives set out in the EU Daughter Directive for NO2: that is, within +25%.

- 68% of the 38 laboratories displayed an average measurement precision of 2.54 ppb.

- However, the range in bias of measurement data was large: -39% to +58%.

- The average precision associated with the measurement data reported by the UK NO2 Network was estimated at 2.54 ppb.

- The recommended limits of acceptable bias were established at 25%. A recommended target for precision of 2.5 ppb was suggested.

- Poor extraction efficiency during analysis was identified as the most likely contributory factor resulting in under estimation by diffusion tubes relative to the chemiluminescent technique and variability in laboratory performance.

- Laboratories showing a negative bias (under estimating) were recommended to investigate their procedures for tube preparation and nitrite extraction from the tube. They were advised to ensure that excess triethanolamine (TEA) was loaded on the absorbent grid of the diffusion tubes and also that all absorbed nitrite was made available for analysis by the extraction phase.

- It was predicted that improved extraction and sampling efficiency may result in average bias in diffusion tube measurements relative to the EU prescribed chemiluminescence method tending to the positive. Also, the reduction in the overall variability in analyses may also be expected, thus improving the comparability of measurements in the UK NO2 Network.

### 1.3 Scope and Objectives of 1999 Field Intercomparison

The 1999 intercomparison was intended to build upon the experience of the 1998 study, and further improve performance of laboratories participating in the UK NO2 Network. The objectives of the 1999 field intercomparison exercise were as follows:

- To estimate bias and precision, under normal field operating conditions, for all laboratories performing analysis in the UK NO2 Network during 1999.

- To compare the estimated bias and precision in 1999 with the results from the field intercomparisons in 1998 and earlier years.

- To investigate whether the method of tube preparation introduces systematic differences in performance of diffusion tubes, for the two main methods.

- To investigate how techniques for preparing, exposing and analysing diffusion tubes could be improved to reduce uncertainty.

## 2. Organisation of the 1999 Field Intercomparison

Thirty-three analytical laboratories providing diffusion tube analysis for UK NO2 Network participants were invited to take part in the 1999 field intercomparison exercise.

### 2.1 Exposure Details

Seven nitrogen dioxide diffusion tubes (six exposure tubes and one travel blank) were supplied by each laboratory, for exposure over a period of one month, August 1999. As in the 1998 study, diffusion tubes were exposed simultaneously, upon purpose made exposure racks located immediately adjacent to the automatic chemiluminescent NOx monitoring equipment installed at the DETR’s Automatic Urban Network (AUN) site at Walsall Alumwell. The exposure time was a 28-day period from 4th August to 1st September 1999.

Upon completion of exposure, diffusion tubes were capped and the exposure time noted. The exposed samplers and travel blanks were then returned to the supplying laboratory for analysis. Travel blanks accompanied exposure tubes to and from the test site. They were isolated in sealed sample bags, and refrigerated throughout the exposure period.

The participating laboratories sent their analytical results to AEA Technology for collation. Results reported in micrograms per cubic metre (µg m-3) were converted to parts per billion (ppb) using a conversion factor of 0.523.

### 2.2 Statistical processing of measurement data

Prior to the interpretation of the measurement data supplied by the laboratories, outlying data within the datasets for each laboratory were removed using the Grubb’s Test. This statistical test was used in an iterative process, at a probability level of P=0.05. It was assumed that the sample distribution was normal. In order to enable a comparison of 1999 data with similar data from the 1998 field intercomparison, outliers from the 1998 dataset were removed retrospectively using the same test.

The decision to remove outliers was taken in order to provide better average estimates of laboratory performance and also take into account the levels of screening employed within the UK NO2 Network whereby abnormally low or high results are investigated and deleted from the Network’s dataset where appropriate. The presence of this level of screening was not recognised in the summary report for the 1998 intercomparison exercise.

## 3. Summary of Results from the Field Testing Exercise

### 3.1 Comparison with chemiluminescent technique

The average diffusion tube measurement for the period 4th August - 1st September 1999 was 17.7ppb, with a standard deviation of 2.1ppb. The average measurement for the same period, from the co-located chemiluminescent NOx analyser, was 19ppb. Table 1 presents the bias as a percentage for the averaged measurement by a laboratory relative to the average measurement from the automatic chemiluminescent analyser for the same period. Code numbers are used to identify each laboratory. Table 1 shows that only three laboratories exhibited a bias relative to the automatic analyser of greater than the target of 25%, and of these three none exceeded this target by more than 6%.

Table 1 Average bias and standard deviation of NO2 diffusion tube measurements, by laboratory

| Laboratory code | % Bias relative to automatic analyser 1999 | Standard Deviation 1999 | % Bias relative to average diffusion tube conc. 1999 | % Bias relative to average diffusion tube conc. 1998 |
| --- | --- | --- | --- | --- |
| 1 |  |  |  | +49.0 |
| 2 | +4.9 | 2.76 | +12.5 | +16.8 |
| 5 | -10.6 | 2.97 | -4.2 | -21.7 |
| 6 | -4.3 | 0.28 | +2.6 | +9.4 |
| 7 | -19.9 | 1.67 | -14.2 | -1.8 |
| 8 | -21.5 | 1.69 | -15.9 | -26.2 |
| 9 | -14.6 | 3.51 | -8.4 | -16.5 |
| 10 | -5.9 | 5.09 | +0.9 | -0.5 |
| 11 |  |  |  | -23.6 |
| 12 | -18.2 | 1.70 | -12.3 | -5.9 |
| 13 | +17.0 | 1.35 | +25.4 | +40.2 |
| 15 | +4.2 | 1.84 | +11.7 | +26.7 |
| 16 | +9.6 | 0.68 | +17.4 | +11.8 |
| 17 | -25.2 | 2.74 | -19.8 | +7.6 |
| 18 | +5.2 | 1.37 | +12.7 | +16.0 |
| 19 | +11.1 | 0.90 | +19.1 | -26.3 |
| 21 | -30.7 | 0.84 | -25.7 | +12.8 |
| 22 | -22.9 | 2.35 | -17.4 | -23.5 |
| 23 |  |  |  | +57.9 |
| 26 | -3.0 | 1.67 | +4.0 | +10.3 |
| 27 | -0.8 | 2.44 | +6.4 | +18.1 |
| 28 | -28.7 | 2.83 | -23.5 | -16.6 |
| 30 | +10.1 | 1.60 | +18.0 | +22.0 |
| 32 | +14.7 | 3.69 | +22.9 | +3.2 |
| 34 | -5.5 | 3.20 | +1.3 | -11.1 |
| 35 | -21.7 | 1.19 | -16.1 | -7.8 |
| 36 | -2.3 | 1.89 | +4.7 | +21.0 |
| 37 | -10.0 | 0.50 | -3.5 | +19.9 |
| 38 | +7.9 | 3.23 | +15.6 | -38.2 |
| 39 |  |  |  | +21.9 |
| 40 | +4.2 | 2.58 | +11.7 | -1.7 |
| 41 | -23.7 | 1.25 | -18.2 | -18.2 |
| 42 | -15.3 | 1.14 | -9.2 | -39.3 |
| 43 |  |  |  | -23.2 |
| 44 |  |  |  | +1.6 |
| 47 | -19.9 | 3.13 | -14.1 | -13.8 |
| 48 | -17.0 | 3.05 | -11.0 | -8.5 |
| 49 | +24.4 | 2.26 | +33.4 | +2.6 |
| 50 | -12.8 | 2.90 | -6.6 |  |
| Average | -6.7 | 2.13 | 0.0 | 1.2 |

Reference concentration (chemiluminescent automatic analyser) = 19 ppb

Average diffusion tube measurement = 17.7 ppb

Table 1 also indicates that on average laboratory analyses underestimated by approximately 7% relative to the automatic analyser. This observation is presented graphically by Figure 1, which shows the distribution of bias in averaged laboratory measurements relative to the chemiluminescent analyser. Clearly, the majority of analyses under estimate concentrations relative to the automatic technique. It is also clear that a large proportion of the laboratories exhibiting under reads use diffusion tubes prepared by the same method. This observation will be further discussed in section 4.

The 1998 field intercomparison also identified an average negative bias relative to the automatic method. As explained in the 1998 intercomparison report, historically, diffusion tubes have been shown to over-read, rather than under-read relative to chemiluminescent measurements by up to 30% (Campbell et al)6. Over-read in diffusion tubes may be attributed to the individual and combined effect of three interfering factors:

- the shortening of the diffusive path length by wind (Atkins et al)7,

- blocking of UV light resulting in reduced NO2 photolysis in the tube (Heal et al)8 and

- the effects of PAN (Atkins et al)7.

The only known interference effect which might reduce the amount of NO2 absorbed on to the diffusion tube, is a possible photochemical reaction between the TEA-nitrite complex with light. This effect, however, has been largely minimised by the widespread use of opaque diffusive end caps (Stevenson et al).

In the 1998 field intercomparison, three tubes from each batch were analysed by a standard laboratory; these tubes, in general, showed less negative bias than those analysed by the supplying laboratory. It was concluded that the standard laboratory was achieving better nitrite extraction from exposed diffusion tubes during analysis. As a result, extraction efficiency was identified as a major source of analytical uncertainty in diffusion tubes. The report recommended that all laboratories optimise their extraction techniques, and ensure that a sufficient excess of absorbent is available throughout exposure. On the basis of the data presented on the 1999 field intercomparison, a problem with variable nitrate extraction is still prevalent in a number of laboratories. The recommendation to optimise extraction techniques therefore remains.

### 3.2 Bias relative to the average diffusion tube measurements

By way of a comparison of laboratory performance in previous intercomparison exercises, Table 1 also presents the average bias in measurements relative to the overall average of all diffusion tube measurements for both 1998 and 1999 field exercises. Outliers have been identified and removed from both datasets using Grubb's test. From Table 1, the overall performance of individual laboratories can be assessed. During the 1999 field intercomparison exercise, measurements from 30 of the 33 laboratories (91%) are within 25% of the average diffusion tube measurement. This represents an improvement relative to the 1998 study, where only 79% had an average bias within 25%. The range in the average bias exhibited by laboratories in 1999 was similarly less than that identified in 1998; -30.7% to +24.4% in 1999, as opposed to -39% to +58% in 1998. It should be noted, however, that the observed range in bias during 1999 remains unacceptably large. Clearly, further efforts are required in order to reduce this variability to an acceptable level.

The overall distribution of bias in measurements, (relative to the mean of all the measurements) is presented in Figure 2. This plot shows that the distribution bias in the data for each laboratory for August 1999 was close to normal. None of the laboratories show evidence of extreme outlying performance.

## 4. investigation of systematic differences in diffusion tube performance due to preparation technique

Currently, there are two main preparation techniques used in the production of NO2 diffusion tubes used in the UK NO2 Network. The diffusion tubes may be prepared using:

- Triethanolamine (TEA) in a 50:50 v/v solution with acetone, or

- Triethanolamine (TEA) in a 50:50 v/v solution with water and a proprietary surfactant (Brij).

From the data obtained in the 1998 field intercomparison trial, it was identified that the two preparation techniques may produce diffusion tubes with different performance characteristics and a requirement was placed to further investigate the effect of preparation technique upon diffusion tube performance.

Figure 1 presented the average bias in the data for each participating laboratory relative to the chemiluminescent analyser. Also indicated is the preparation technique for the diffusion tubes used by each laboratory; light columns indicate diffusion tubes prepared using 50% v/v TEA in water with surfactant, dark columns indicate diffusion tubes prepared using 50% v/v TEA in acetone. The majority of the diffusion tubes (71%) exhibiting a negative bias were prepared using TEA in water, while the majority of those exhibiting a positive bias (73%) were prepared using TEA and acetone. This observation further exemplifies the potential effect of diffusion tube preparation on sampling performance. Indeed, the evidence would be fairly categorical if it were not for the results from a small number of laboratories, which show atypical diffusion tube performance for the preparation technique used.

In order to investigate further the potential effect of preparation technique upon diffusion tube performance, an additional nine tubes of each preparation technique were exposed as part of an independent test. Three diffusion tubes of each type were retained as travel blanks. After exposure, both sets of samplers were analysed by the AEA Technology laboratory.  The results are shown in Table 2.

Table 2.  Comparison of diffusion tube measurements  prepared using 50% v/v TEA/acetone and 50% v/v TEA/Water+ Brij surfactant. Analysis by AEA Technology.

| Tube | Mass Nitrite (µg) | NO2 ppb | % Bias |  | Tube | Mass Nitrite (µg) | NO2 ppb | % Bias |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| TEA/Acetone 1 | 1.70 | 18.6 | -2 |  | TEA/Water 1 | 1.19 | 12.6 | -34 |
| TEA/Acetone 2 | 1.92 | 21.0 | +11 |  | TEA/Water 2 | 1.16 | 12.2 | -36 |
| TEA/Acetone 3 | 1.85 | 20.3 | +7 |  | TEA/Water 3 | 1.54 | 16.2 | -15 |
| TEA/Acetone 4 | 1.87 | 20.5 | +8 |  | TEA/Water 4 | 1.85 | 19.5 | +2 |
| TEA/Acetone 5 | 1.84 | 20.2 | +6 |  | TEA/Water 5 | 1.83 | 19.4 | +2 |
| TEA/Acetone 6 | 2.00 | 21.9 | +15 |  | TEA/Water 6 | 1.16 | 12.2 | -36 |
| TEA/Acetone 7 | 2.11 | 23.2 | +22 |  | TEA/Water 7 | 1.16 | 12.2 | -36 |
| TEA/Acetone 8 | 2.06 | 22.0 | +16 |  | TEA/Water 8 | 1.59 | 16.8 | -12 |
| TEA/Acetone 9 | 1.95 | 21.4 | +13 |  | TEA/Water 9 | 1.04 | 11.0 | -42 |
| TEA/Acetone Blank 1 | 0.02 | - | - |  | TEA/Water Blank 1 | 0.03 | - | - |
| TEA/Acetone Blank 2 | 0.02 | - | - |  | TEA/Water Blank 2 | 0.03 | - | - |
| TEA/Acetone Blank 3 | 0.02 | - | - |  | TEA/Water Blank 3 | 0.02 | - | - |
| Mean |  | 21.0 | +11 |  |  |  | 14.7 | -23 |
| Std Dev. |  | 1.31 |  |  |  |  | 3.31 |  |

Table 2 clearly shows that the diffusion tubes prepared using acetone as the solvent for TEA exhibited a positive bias relative to the chemiluminescent analyser, (on average +11%), while those prepared using water and surfactant as the solvent, exhibited negative bias, ( on average -23%). The standard deviation of the results from diffusion tubes prepared using the TEA in water method is approximately 2.5 times that of the results from diffusion tubes prepared using TEA in acetone.

The results presented in Table 2 are broadly consistent with the results obtained by the supplying laboratory in the main 1999 intercomparison trial (presented in Table 3 below). However, in the main trial both supplying laboratories achieved similar standard deviations.

Table 3 Comparison of diffusion tube measurements prepared using 50% v/v TEA/acetone and 50% v/v TEA/Water+ Brij surfactant, analyses performed by supplying laboratories

| Laboratory 1 |  |  |  |  | Laboratory 2 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Tube | Mass Nitrite (µg) | NO2 ppb | % Bias |  | Tube | Mass Nitrite (µg) | NO2 ppb | % Bias |
| TEA/Acetone 1 | 1.99 | 21.9 | +15 |  | TEA/Water 1 | 1.57 | 16.4 | -14 |
| TEA/Acetone 2 | 1.89 | 20.8 | +9 |  | TEA/Water 2 | 1.36 | 14.3 | -25 |
| TEA/Acetone 3 | 1.50 | 16.4 | -14 |  | TEA/Water 3 | 1.30 | 13.6 | -28 |
| TEA/Acetone 4 | 1.80 | 19.8 | +4 |  | TEA/Water 4 | 1.68 | 17.7 | -7 |
| TEA/Acetone 5 | 1.79 | 19.8 | +4 |  | TEA/Water 5 | 1.36 | 14.3 | -25 |
| TEA/Acetone 6 | 1.85 | 20.3 | +7 |  | TEA/Water 6 | 1.62 | 17.0 | -10 |
| TEA/Acetone Blank | 0.02 | - | - |  | TEA/ Blank | 0.03 | - | - |
| Average |  | 19.8 | +4 |  |  |  | 15.5 | -18 |
| Std. Dev. |  | 1.83 |  |  |  |  | 1.70 |  |

## 5. Recommended Guidelines for Laboratory Performance

### 5.1 Data Quality Objectives for the Network

Under the European Union Daughter Directive for NO2,9 data quality objectives have been set out for the overall accuracy of indicative monitoring techniques (i.e. diffusive monitoring). In the case of diffusion tube monitoring, annual average NO2 concentration data must comply with these data quality objectives to enable comparison of measurement data with the Daughter Directive for NO2. For indicative monitoring the data quality objective has been set at 25%. Hence, it is recommended that on average, diffusion tube measurements should be within 25% of the reference concentration.

This objective has been used within the UK NO2 Network since 1997 as the data quality objective for the screening of diffusion tube data with unsatisfactory quality. It is broadly consistent with indicators of good laboratory performance used within the WASP proficiency testing scheme for NO2 diffusion tubes, and also the UK NO2 Network Laboratory Performance Testing Scheme used from 1996-1999. Applying this criteria to the 1999 field intercomparison exercise, laboratories that are, on average, within approximately 25% of the reference value, may be recognised as performing satisfactorily. Conversely, laboratories with an average bias significantly greater than 25% will have performed unsatisfactorily.

It must be noted however, that the precision of analytical measurements should also be taken into account, as it is possible to achieve an average bias of less than 25% with very imprecise measurements (i.e. purely by chance). Last year, the average precision achieved by all laboratories in the 1998 field intercomparison was 2.5ppb. This value was recommended as a "target" for the precision of diffusion tube measurements. However, as around half of the laboratories will inevitably have precision values greater than the average, this was a rather stringent target. Therefore, in this 1999 study we have taken a figure of 3ppb as an arbitrary guideline for acceptable precision. The average precision achieved in the 1999 study was 2.1ppb.

Using these guidelines, three of the 33 laboratories taking part in this field intercomparison exercise (9%) produced measurement data with average bias >25%, relative to the reference measurement from the chemiluminescent analyser. (In all three cases the bias was negative). Seven laboratories (21%) showed an average precision greater than our arbitrary guideline of 3 ppb. No laboratories produced measurement data with an average bias >25% and an average precision >3ppb.

### 5.2 Data quality objectives for the Air Quality Strategy (AQS)

Under the AQS, local authorities are obliged to assess and review the air quality within their authority10. Diffusion tube surveys may be used as screening tools within these assessments, as the annual average data derived from NO2 diffusion tubes may be directly comparable with the AQS air quality objective for the end of 2005 (21 ppb for annual mean NO2).

In recognition of the evidence for a potential bias in diffusion tube measurement data, and in the absence of a methodology for estimating accuracy as defined by EC Directive 1999/30/EC, information on the bias (relative to the chemiluminescent technique) and precision of diffusion tube measurement data must be presented for the period of monitoring. Where appropriate at Stage 3, scaling factors may also be applied to the diffusion tube measurement data to correct for any systematic bias. Any use of scaling factors must be reported.

In addition to the data presented in this report, performance testing data from the UK NO2 Network Laboratory Performance Testing Scheme 1996-1999 and the new WASP proficiency scheme are available from laboratories and may be used to provide a further information on the general performance of laboratories.

## 6. Comparison with Other Field Intercomparison Exercises

A comparison of the average performance data produced from the 1999 field intercomparison exercise with data from exercises performed in 1994, 1995 and 1998 indicates that the performance characteristics of the laboratories have improved. Table 4 presents the average performance statistics from intercomparisons between 1994 and 1999.

Table 4. Average Laboratory Performance in the UK NO2 Network Field Intercomparison Exercises 1994-1999

|  | Average Performance in UK NO2 Network Field Intercomparisons |  |  |  |
| --- | --- | --- | --- | --- |
|  | 1994 | 1995 | 1998 | 1999 |
| Average Bias (%) | -11 | 2 | 1.7 | -6.7 |
| Maximum Bias (%) | 118 | 118 | 58 | 24 |
| Minimum Bias (%) | -96 | -87 | -39 | -31 |
| Standard Deviation of Bias | 39 | 39 | 22 | 15 |
| Precision (ppb) | 2.04 | 0.83 | 2.54 | 2.13 |

Table 4 shows that the performance of participating laboratories has generally improved since 1994. The average bias has increased in magnitude compared with 1998, to -6.7%, however, this may reflect the reduction in the number of laboratories exhibiting large positive bias. The range of bias has reduced, and is illustrated by the gradual reduction in maximum, minimum, and standard deviation in laboratory bias, presented in Table 4.

It is also possible to compare the bias in measurements, relative to the average diffusion tube measurement for individual laboratories in the 1998 and 1999 intercomparisons. This is illustrated in Figure 3. Grubb's test has been used to identify outliers in both 1998 and 1999 datasets. Thirty-two laboratories took part in both consecutive years' intercomparisons. The majority (56%) of laboratories exhibited a smaller percentage bias (relative to the average diffusion tube measurement) in 1999 than in 1998.

## 7. Conclusions

The following conclusions may be drawn from this field intercomparison exercise:

- 30 of the 33 laboratories (91%) performed to the data quality objectives set out in the proposed EU Daughter Directive for NO2, that is, bias within 25%.

- 26 laboratories (79%) displayed an average measurement precision within our arbitrary guideline of 3 ppb.

- The range in bias of measurement data, relative to the average of all the diffusion tube measurements in the study, was -30.7% to +24.4%. This represents an improvement relative to the 1998 intercomparison in which the range in bias was -39% to +58%.

- The average bias associated with the national monthly and annual average measurement data reported by the UK NO2 Network is estimated at - 6.8% relative to the chemiluminescent technique. This is a larger negative bias than that obtained in the 1998 intercomparison and may reflect the reduced occurrence of large positive bias, while large negative bias continues to be observed.

- The average precision associated with the measurements in this study was 2.1 ppb.

- Local authorities are recommended to ensure that the performance for their analytical laboratory is within approximately 25% (for bias) and 3ppb (for precision). It is also recommended that additional information on performance testing and validation exercises be utilised in assessing laboratory performance where this is available. Full details on the performance of individual laboratories are available direct from the laboratory.

- Further evidence of the effect of diffusion tube preparation technique upon the performance of diffusion tubes was identified. Tubes prepared using the TEA in acetone method typically exhibited a positive bias, while tubes prepared using the TEA in water method typically exhibited a negative bias.

## 8. recommendations for further improvement

- The following recommendations are proposed to improve the quality data within the UK NO2 Network and for further work to better define potential uncertainties:

- Laboratories displaying a negative bias are strongly encouraged to review their procedures for preparing diffusion tubes and extracting nitrite after exposure; the most likely cause of under-read is incomplete extraction of absorbed NO2 into the analysed sample and/or insufficient loading of the TEA absorbent onto the diffusion tube.

- The extraction efficiency of a laboratory may be effectively tested and monitored on a monthly basis using the WASP programme for NO2 diffusion tubes. As low extraction efficiency is likely to be the largest individual source of negative bias in diffusion tube measurements, it is strongly recommended that checks are made to ensure that laboratory performance in the WASP scheme is within ±25% of the target concentration.

- Laboratory based tests investigating the effect of diffusion tube preparation technique should be performed, with a view to standardisation of the preparation techniques used within the UK NO2 Network.

## 9. acknowledgements

AEA Technology would like to thank all the laboratories who participated in this study for their co-operation.

## 10. references

- Bush, T, Mooney D and Stevenson, K.J. UK Nitrogen Dioxide Survey 1996. AEA Technology, National Environmental Technology Centre. AEAT - 2279, ISBN 0-7058-1760-1.

- Stevenson, K.J. and Bush, T. UK Nitrogen Dioxide Survey Results for the First Year-1993. AEA Technology, National Environmental Technology Centre, AEA/CS/RAMP/16419031/002, 1994.

- Stevenson, K.J. and Bush, T. UK Nitrogen Dioxide Survey 1994. AEA Technology, National Environmental Technology Centre. AEAT - 0085, ISBN 0-7958-1711-3.

- Stevenson, K.J., Bush, T. and Mooney D. UK Nitrogen Dioxide Survey 1995. AEA Technology, National Environmental Technology Centre. AEAT - 0912, ISBN 0-7058-1703-X.

- Bush, T. Summary Results from the UK NO2 Network Field Intercomparison Exercise 1998. Available from DETR Air Quality Archive web site  http://www.aeat.co.uk/netcen/airqual/reports/home.html or from AEA Technology.

- Campbell G. W, Stedman J. R, and Stevenson, K.J. A Survey of Nitrogen Dioxide Concentrations in the United Kingdom Using Diffusion Tubes, July - December 1991. Atmospheric Environment Vol. 28, No. 3, pp 477-486, 1994.

- Atkins C.H.F, Sandalls J, Law D.V, Hough A.M, and Stevenson K.J. The Measurement of Nitrogen Dioxide in the Outdoor Environment Using Passive Diffusion Tube Samplers. Harwell Energy & Environment Report. No. AERE-R 12133. February 1986.

- Heal M. R, O’Donoghue M. A, and Cape J. N. Overestimation of urban nitrogen dioxide by passive diffusion tubes: a comparative exposure and model study. Atmospheric Environment 33, pp 513-524, 1999.

- Council Directive EC 1999/30/EEC relating to limit values for sulphur dioxide, nitrogen dioxide and oxides of nitrogen, particulate matter and lead in ambient air. 22 April 1999.

- LAQM.TG1(98), Monitoring for air quality review and assessments. Part IV The Environment Act 1995, Local Air Quality Management, April 1998. Department of the Environment Transport and the Regions, Scottish Office and Welsh Office. HMSO, ISBN 0 11 753429 3.